## Slide 1: BioRobotics

## Slide 2
Why EMG?
Prosthetics
Intuitive control of robotic limbs using muscle signals
Gesture Control
Hands-free interaction with devices and computers
Rehabilitation
Biofeedback for stroke recovery and physical therapy
Exoskeletons
Assistive devices that respond to user intent
EMG provides a direct window into motor intent

## Slide 3
The Big Picture
Brain
Intent
→
Muscle
Contraction
→
EMG
Signal
→
Signal
Processing
→
Classification
→
Robot
Control

## Slide 4
The Motor Unit
The fundamental unit of muscle control
Motor neuron + all muscle fibers it innervates
Each motor unit fires as a single unit (all-or-none)
Force control: recruit more motor units or increase firing rate

## Slide 5
What is EMG?
Electromyography (EMG)
Electrical recording of muscle activity during contraction
Frequency Range
2 - 500 Hz
(most energy 20-150 Hz)
Amplitude
50 µV - 5 mV
(depends on electrode)
Electrode Types
Surface: non-invasive
Intramuscular: high resolution
We'll focus on surface EMG (sEMG)

## Slide 6
From Action Potential to Contraction
Action potential: 1-3 ms electrical spike
Twitch: 10-100 ms mechanical response
Tetanus: fused contractions at high firing rates (>30 Hz)

## Slide 7
EMG as Superposition

## Slide 8
EMG Signal

## Slide 9
Electrode Configuration
Bipolar differential: 2 electrodes + reference
Measures voltage difference to reject common noise
Placement matters: along muscle fiber direction
Inter-electrode distance: 1-2 cm typical
Myo Armband Specifications:
8 channels (dry electrodes)
200 Hz sampling rate
8-bit resolution
Wireless via Bluetooth
Good for gesture classification

## Slide 10
Noise Sources in EMG

## Slide 11
Filtering Fundamentals
Nyquist Theorem
Sample at ≥ 2× max frequency to avoid aliasing
Myo: 200 Hz → can capture up to 100 Hz
Common Filters
High-pass (20 Hz): remove motion artifacts
Low-pass (450 Hz): anti-aliasing
Notch (50/60 Hz): powerline rejection
Butterworth: smooth frequency response

## Slide 12
Signal Processing Pipeline
Raw → Bandpass → Rectify → Smooth
Each step reveals different information

## Slide 13
Envelope Extraction
Rectification
y(t) = |x(t)|
Convert bipolar signal to unipolar
Makes averaging meaningful
Linear Envelope
Low-pass filter the rectified signal
Smooths rapid fluctuations
Tracks muscle activation level
Used for proportional control
Proportional control: Envelope amplitude → Robot speed/force

## Slide 14
Time-Domain Features (Hudgins Set)
MAV: Mean Absolute Value - overall activation level
RMS: Root Mean Square - signal power
ZC: Zero Crossings - frequency content proxy
WL: Waveform Length - signal complexity
SSC: Slope Sign Changes - high-frequency content

## Slide 15
Feature | Description | Physiological Correlation | Application
Median Frequency (MDF) | Frequency dividing power spectrum into equal halves | ↓ Decreases with FATIGUE (spectral shift to lower freq) | Fatigue monitoring, ergonomic assessment
Mean Frequency (MNF) | Centroid of the power spectral density | ↓ Decreases with FATIGUE (tracks conduction velocity) | Endurance studies, rehabilitation progress
RMS / MAV | Root mean square / Mean absolute value | ↑ Increases with FORCE (motor unit recruitment) | Proportional control, force estimation
Zero Crossings (ZC) | Number of times signal crosses zero per second | ~ Reflects FIRING RATE (motor unit discharge) | Gesture classification, MUAP analysis
Waveform Length (WL) | Cumulative length of the EMG waveform | ↑ Increases with ACTIVITY (amplitude + complexity) | Onset detection, activity quantification
💡 Key Insight: Median frequency (MDF) is the gold standard for fatigue detection. As muscle fatigues, conduction velocity decreases and action potentials broaden, shifting power to lower frequencies. A 10-15% drop in MDF typically indicates significant fatigue.

## Slide 16
Evolution of EMG Classification
1960s-80s: Simple threshold detectors
1990s-2010s: Pattern recognition (LDA dominates)
2010s+: Deep learning approaches

## Slide 17
Linear Discriminant Analysis (LDA)
The Workhorse of EMG Classification
Projects features to maximize class separation
Assumes Gaussian distributions, shared covariance
Fast training, fast inference, interpretable
QDA: allows different covariance per class
Typical Pipeline
8 EMG channels
→ Window (200ms)
→ Features per channel
→ Concatenate (8×4=32 features)
→ LDA classifier
Real-world accuracy: 90-98% for 6-10 gestures in controlled conditions

## Slide 18
Deep Learning Approaches
CNNs: Learn spatial patterns across channels
1D Convolutions: Temporal feature learning
RNNs/LSTMs: Capture temporal dynamics
Transformers: Emerging approach
Advantages:
Learn features automatically
Better generalization (sometimes)
Challenges:
Need more training data
Harder to interpret
Computational requirements

## Slide 19
The Robustness Challenge
Lab accuracy ≠ Real-world performance
Electrode Shift
Small position changes degrade accuracy
Muscle Fatigue
Signal characteristics change over time
Cross-Session
Models don't transfer between sessions
Arm Position
Different limb positions → different signals
Active research area: adaptive algorithms, domain adaptation, user calibration

## Slide 20
Prosthetic Control
Evolution of Prosthetic EMG
Direct Control (1960s): 1 electrode → 1 function
Pattern Recognition: Multiple gestures from muscle patterns
Simultaneous Control: Multiple DOF at once
Commercial Systems
• LUKE Arm (Mobius Bionics)
• bebionic (Ottobock)
• COAPT pattern recognition
Key insight: Pattern recognition enables more natural, intuitive control

## Slide 21
Rehabilitation & Therapy
Stroke Recovery: EMG biofeedback helps relearn movements
Motor Learning: Real-time feedback accelerates skill acquisition
Gamification: Engaging exercises improve compliance
Biofeedback Loop
Attempt movement
→ EMG detected
→ Visual/audio feedback
→ Reinforces correct patterns

## Slide 22
Exoskeletons & Cobots
Assistive Exoskeletons: EMG triggers and modulates assistance
Collaborative Robots: EMG for intent detection and safety
Industrial Augmentation: Reduce worker fatigue and injury
Key Advantage
EMG provides intent ~50-100ms before motion begins

## Slide 23
Consumer & Emerging Applications
Myo Armband (2014-2018): First mainstream consumer EMG. Gesture control for presentations, drones, gaming.
Meta Neural Interface: Wristband with high-density EMG for AR/VR. Acquired CTRL-labs (2019).
Gaming & VR: Hands-free control, accessibility
Emerging Applications:
Silent speech (facial EMG)
Music performance (accessibility)
Fitness tracking (muscle activation)
Accessibility (computer control)

## Slide 24
Research Frontiers
High-Density EMG (HD-EMG): 64-256 electrode grids for spatial mapping, motor unit decomposition
Implantable Sensors: Targeted muscle reinnervation (TMR), intramuscular electrodes
Hybrid Brain-Muscle Interfaces: EEG + EMG fusion for intent detection
Emerging Research:
Transfer learning: Adapt to new users
Unsupervised adaptation: Self-updating
Regression control: Continuous force
Real-time edge AI: On-device inference
Open challenge: Robust control that works all day, every day