## Slide 1: BioRobotics

## Slide 2
What is Electrooculography?
EOG is the technique of measuring the corneo-retinal standing potential that exists between the front and back of the eye.
Key Concept: The Eye as a Battery
Cornea (front):
Positively charged (+)
Retina (back):
Negatively charged (−)
Potential difference:
0.4 to 1.0 mV
When eyes rotate:
Dipole orientation changes
Corneo-Retinal Dipole
+
−
Retina (−) Cornea (+)
3 / 32

## Slide 3
How EOG Detects Eye Movements
Eyes Center
0
0
Equal potential at both electrodes
Eyes Right
−
+
Cornea (+) closer to right electrode
Eyes Left
+
−
Cornea (+) closer to left electrode
The Principle
Electrodes placed around the eyes detect changes in the electric field as the eye rotates. The voltage difference between electrode pairs is proportional to the angular displacement of the eye.
4 / 32

## Slide 4
Types of Eye Movements
Saccades
Speed: 30-500°/s
Duration: 20-200 ms
Rapid, ballistic movements between fixation points
Fixations
Speed: < 1°/s
Duration: 200-600 ms
Stable gaze on a target with micro-movements
Smooth Pursuit
Speed: 10-30°/s
Duration: Variable
Slow tracking of moving objects
Blinks
Speed: N/A
Duration: 100-400 ms
Eyelid closure, creates large VEOG artifact
5 / 32

## Slide 5
EOG Electrode Placement
Standard 2-Channel Bipolar Configuration
Channel 1: Horizontal EOG (HEOG)
Electrode placement:
• Right outer canthus (+)
• Left outer canthus (−)
Detects: Left/Right eye movements
Channel 2: Vertical EOG (VEOG)
Electrode placement:
• Above right eye (+)
• Below right eye (−)
Detects: Up/Down movements + Blinks
Ground Electrode:
Center of forehead or behind ear (mastoid) - provides stable reference potential
Face Diagram (see lab handout for detailed placement)
7 / 32

## Slide 6
HEOG vs VEOG Signal Patterns
Gaze Direction
HEOG Response
VEOG Response
Look Right
Positive (+)
~Baseline
Look Left
Negative (−)
~Baseline
Look Up
~Baseline
Positive (+)
Look Down
~Baseline
Negative (−)
Blink
Small artifact
Large spike (+)
💡 Key Insight: HEOG and VEOG are largely independent. Horizontal movements primarily affect HEOG, vertical movements primarily affect VEOG. Blinks create the largest artifact in VEOG.
8 / 32

## Slide 7
EOG Signal Characteristics
Frequency Range
DC - 30 Hz
Very low frequency content
Amplitude
15-200 μV/°
Per degree of eye rotation
Sample Rate
250-500 Hz
Sufficient for most applications
Signal Type
Quasi-DC
Sustained gaze = sustained voltage
Why This Matters for Processing
The low frequency content of EOG means simple low-pass filtering is very effective. However, DC coupling is essential to capture sustained gaze positions, and baseline drift becomes a significant concern.
10 / 32

## Slide 8
EOG vs EMG: Key Differences
Understanding how these signals differ helps you apply the right processing techniques
Property
EOG
EMG
Signal Source
Corneo-retinal potential
Motor unit action potentials
Frequency Range
DC - 30 Hz
20 - 500 Hz
Amplitude
15-200 μV
50 μV - 5 mV
Coupling
DC (essential)
AC (typical)
Signal Nature
Quasi-DC, position-based
Stochastic, activity-based
Baseline
Drifts significantly
Returns to zero
Key Challenge
Drift & blink artifacts
Crosstalk & fatigue
11 / 32

## Slide 9
Signal Appearance: EOG vs EMG
EOG Signal (Look Right → Center → Left)
Step-like changes with sustained levels
EMG Signal (Muscle Contraction)
Stochastic bursts, returns to baseline
Key Processing Implications
EOG:
Low-pass filter effective; baseline correction critical; feature = amplitude/position
EMG:
Band-pass filter needed; RMS/MAV for envelope; frequency features important
12 / 32

## Slide 10
EOG Hardware Requirements
Biosignal Amplifier
High input impedance, DC-coupled input, 2+ channels
Electrodes
Ag/AgCl disposable or reusable cup electrodes
Conductive Gel
Electrolyte gel for low impedance contact
ADC
12-16 bit resolution, 250-500 Hz sample rate
Software
Real-time acquisition and processing capability
Typical Specifications
Sample Rate
250-500 Hz
Resolution
12-16 bit
Input Range
±2-5 mV
CMRR
>80 dB
Bandwidth
DC-100 Hz
⚠️ Critical: DC Coupling Required
Unlike EMG, EOG requires DC coupling to capture sustained gaze positions. AC coupling will filter out the steady-state voltages that indicate where the eye is looking. Ensure your amplifier supports DC mode or has a very low high-pass cutoff (<0.1 Hz).
21 / 32

## Slide 11
Electrode Placement Tips
Skin Preparation
• Clean with alcohol wipe
• Allow to dry completely
• Light abrasion if needed
• Check impedance <10kΩ
Electrode Placement
• Outer canthi for HEOG
• Above/below eye for VEOG
• ~1 cm from eye
• Symmetrical positioning
Common Mistakes
• Electrodes too far from eye
• Asymmetric placement
• Poor skin contact
• Dried out gel
Quality Checks
• Look L/R - verify HEOG
• Look U/D - verify VEOG
• Blink - large VEOG spike
• Stable baseline at rest
22 / 32

## Slide 12
Signal Quality Considerations
Factor
Good
Acceptable
Poor
Electrode Impedance
<5 kΩ
<10 kΩ
>20 kΩ
Baseline Stability
Minimal drift
Slow drift (<50μV/min)
Rapid drift
Signal-to-Noise
>10:1
>5:1
<3:1
60Hz Noise
Not visible
Minor ripple
Dominant
Quick Troubleshooting
High impedance:
Re-prep skin, add gel, check cable connections
Baseline drift:
Allow electrodes to stabilize (2-3 min), check gel moisture
60Hz noise:
Check ground electrode, shield cables, move away from power sources
23 / 32

## Slide 13
Common EOG Artifacts
Blink Artifacts
Channel: VEOG primarily
Large positive spikes (100-400μV)
Detect and reject, or use as control signal
Handling:
Detect and reject, or use as control signal
Muscle Artifacts
Channel: Both channels
High-frequency noise bursts
Low-pass filter (<30 Hz)
Handling:
Low-pass filter (<30 Hz)
Baseline Drift
Channel: Both channels
Slow wandering of baseline
High-pass filter (0.1 Hz) or adaptive baseline
Handling:
High-pass filter (0.1 Hz) or adaptive baseline
Power Line (60Hz)
Channel: Both channels
Regular sinusoidal interference
Notch filter or improve shielding
Handling:
Notch filter or improve shielding
24 / 32

## Slide 14
Blink Detection
Blinks as Artifacts vs. Control Signals
As Artifacts (Remove)
When analyzing gaze patterns:
• Detect using threshold crossing
• Mark affected epochs
• Reject or interpolate
• Typical duration: 100-400ms
As Control Signals (Use)
For human-computer interaction:
• Single blink = click/select
• Double blink = confirm/enter
• Long blink = cancel/back
• Blink patterns as commands
Simple Detection Algorithm
1. Baseline-correct VEOG signal → 2. Set threshold (e.g., 3× std dev) → 3. Find samples above threshold → 4. Check duration (50-400ms typical blink) → 5. Mark as blink or reject
25 / 32

## Slide 15
Feature Extraction for Gaze Classification
Amplitude Features
• Mean amplitude (HEOG, VEOG)
• Peak amplitude
• Peak-to-peak range
Most discriminative for gaze direction
Statistical Features
• Standard deviation
• Skewness
• Kurtosis
Capture signal variability
Temporal Features
• Saccade velocity (max derivative)
• Fixation duration
• Zero crossings
Capture eye movement dynamics
Energy Features
• RMS amplitude
• Signal energy
• HEOG/VEOG energy ratio
Robust to baseline offset
27 / 32

## Slide 16
Classification Approaches
Applying ML methods you've learned to EOG data
Threshold-based
Simple amplitude thresholds for L/R/U/D
✓ Fast, interpretable
⚠ Sensitive to calibration
SVM
Linear or RBF kernel classifiers
✓ Works well with small data
⚠ Requires feature engineering
PCA/ICA
Dimensionality reduction preprocessing
✓ Separates sources
⚠ May lose information
Neural Networks
CNN/RNN on raw or processed signals
✓ Learns features
⚠ Needs more training data
Lab Focus: You'll apply PCA/ICA for dimensionality reduction and SVM for classification of gaze directions.
28 / 32

## Slide 17
EOG Applications
♿
Assistive Technology
Eye-controlled interfaces for motor disabilities
😴
Drowsiness Detection
Driver monitoring and alertness systems
🌙
Sleep Studies
REM stage detection in polysomnography
🥽
VR/AR
Eye tracking for foveated rendering
🏥
Clinical Diagnostics
Nystagmus and vestibular assessment
14 / 32

## Slide 18
Assistive Technology
Eye-Controlled Interfaces for Motor Disabilities
Use Cases
Wheelchair navigation control
Computer cursor control
Communication devices (AAC)
Environmental control (lights, TV)
Robotic arm assistance
Why EOG for Assistive Tech?
✓ Works with closed eyes (unlike cameras)
✓ Works in any lighting condition
✓ Low cost compared to eye trackers
✓ Simple electrode setup
✓ Real-time processing possible
Target Users:
ALS/MND patients, spinal cord injury, locked-in syndrome, severe cerebral palsy, muscular dystrophy — individuals who retain eye movement control when other motor functions are impaired.
15 / 32

## Slide 19
Drowsiness Detection
Driver Monitoring and Alertness Systems
EOG Indicators of Drowsiness
PERCLOS
Percentage of eye closure over time
Blink Duration
Increases with fatigue (>300ms = drowsy)
Blink Frequency
Irregular patterns indicate fatigue
Slow Eye Movements
Rolling movements before microsleep
🚗
Drowsy driving causes
100,000+
crashes per year (US)
EOG can detect drowsiness 1-2 minutes before a microsleep event
Advantage over camera-based systems: EOG works regardless of sunglasses, lighting conditions, or head position. VEOG directly measures eyelid position.
16 / 32

## Slide 20
Sleep Studies
REM Detection in Polysomnography (PSG)
Wake
Blinks, saccades, reading movements
N1 (Light)
Slow rolling eye movements (SEMs)
N2 (Light)
Minimal eye movements
N3 (Deep)
Very rare eye movements
REM
Rapid conjugate eye movements
Why REM Matters
• Memory consolidation
• Dreaming occurs
• Emotional processing
• 20-25% of adult sleep
EOG is essential for accurate REM staging
Conjugate Eye Movements:
Both eyes move together in the same direction. In REM sleep, these are rapid and occur in bursts. The HEOG channel shows the characteristic rapid deflections that give REM its name.
17 / 32

## Slide 21
VR/AR Eye Tracking
Foveated Rendering and Gaze Interaction
Foveated Rendering
Render high resolution only where the user is looking
• Reduces GPU load by 50-70%
• Enables higher quality headsets
• Critical for wireless VR
Gaze-Based Interaction
Use eye position as an input modality
• Menu selection by looking
• Social eye contact in VR
• Attention-aware interfaces
EOG in VR Headsets
While most commercial VR headsets use camera-based eye tracking, EOG electrodes integrated into the headset padding offer advantages: lower latency, works when eyes are closed, and no IR illumination needed. Research prototypes are exploring hybrid approaches.
18 / 32

## Slide 22
Clinical Diagnostics
Nystagmus and Vestibular Assessment
Nystagmus Detection
Involuntary rhythmic eye oscillations indicating neurological or vestibular issues
Tests: Gaze-evoked, positional, optokinetic
Vestibular Function
Assess inner ear and balance system disorders
Tests: Caloric testing, rotational tests
Oculomotor Testing
Evaluate saccades, pursuit, and vergence
Tests: Saccade accuracy, pursuit gain
Neurological Assessment
Detect brainstem and cerebellar lesions
Tests: Smooth pursuit, gaze holding
Clinical Advantage: EOG provides objective, quantifiable eye movement data that can be compared to normative databases for diagnosis.
19 / 32

## Slide 23
Lab 2: What You Will Do
1
Setup & Calibration
Connect electrodes, verify signal quality, calibrate for your eyes
2
Data Collection
Record EOG for 7 gaze classes: center, left, right, up, down, blink, double-blink
3
Signal Processing
Apply baseline correction and filtering to clean the data
4
Feature Extraction
Extract amplitude, statistical, and temporal features
5
Dimensionality Reduction
Apply PCA and ICA to visualize and separate signal components
6
Classification
Train and evaluate SVM classifiers (Linear and RBF kernels)
Deliverables: Collected data files, completed analysis notebook, answers to 15 conceptual questions
29 / 32

## Slide 24
Key Takeaways
1
EOG measures the corneo-retinal potential that changes with eye rotation
2
Low frequency content (DC-30Hz) requires DC coupling and simple filtering
3
2-channel setup: HEOG for horizontal, VEOG for vertical movements + blinks
4
Different from EMG: quasi-DC signal, position-based rather than activity-based
5
Applications range from assistive tech to sleep studies to VR
6
Classification uses amplitude features with SVM or threshold-based methods
30 / 32

## Slide 25
Resources & Further Reading
1. Bulling, A., et al. (2011). Eye Movement Analysis for Activity Recognition Using Electrooculography. IEEE TPAMI, 33(4).
2. Barea, R., et al. (2002). System for assisted mobility using eye movements based on EOG. IEEE Trans Neural Syst Rehabil Eng.
3. Malmivuo, J. & Plonsey, R. (1995). Bioelectromagnetism. Chapter 28: Electric Signals of the Eye.
4. Iber, C. (2007). The AASM Manual for the Scoring of Sleep and Associated Events. (Sleep staging with EOG)
5. Lab Streaming Layer (LSL) documentation: labstreaminglayer.org
Lab Materials: github.com/AHRTLab/BioRobotics-Lab2
32 / 32